Standard Operating Procedure

| Document Number: |  | Revision: | 2 |
| --- | --- | --- | --- |
| Title: | Time-Course Bioreactor Sampling and Viability Analysis |  |  |
| Effective Date: | May 1, 2026 |  |  |

Trellis Bioworks    Project: CHO-K1 Fed-Batch Optimization    Generated: May 17, 2026

1.0 Purpose

2.0 Scope

3.0 Procedure: Time-Course Bioreactor Sampling and Viability Analysis

CRITICAL REQUIREMENT: All sampling MUST occur within ±5 min of the nominal time point. Deviations >5 min require an in-process deviation (IPD) report before the next time point.

3.1  Hour 0 (T=0) — Inoculation Sampling

Reference time point. Sample drawn immediately after inoculation establishes baseline viability and metabolite concentration.

Figure 1: Inoculation sampling port (post-CIP). Annotated photograph of the bioreactor sampling port showing the sterile septum and integrated luer fitting.

| Time Target | Action | Expected Output / Log |
| --- | --- | --- |
| T=0 min | Aseptically draw 5 mL via sterile sampling port into pre-labelled 15 mL conical. | Sample S-0 logged in eBR; cap torque verified. |
| T+10 min | Perform trypan-blue viability count (Cellometer K2, 1:1 dilution). | Viable cell density (cells/mL); viability %. |
| T+15 min | Submit retained aliquot to in-process QC (pH, glucose). | pH, glucose (g/L) recorded; sample chain-of-custody signed. |

3.2  Hour 3 (T=3) — Foam  Growth Check

Inspection only — no draw unless foam height exceeds the documented limit.

Figure 2: Acceptable vs. excessive foam. Side-by-side comparison of acceptable headspace foam (10 mm) and the rejection threshold.

| Time Target | Action | Expected Output / Log |
| --- | --- | --- |
| T=3 h | Visually inspect headspace foam height; compare to Figure 2. | PASS / FAIL noted in eBR; foam height (mm). |
| T=3 h +5 min | If FAIL: add 0.5 mL sterile antifoam-C via dedicated addition port and re-inspect. | Antifoam lot, volume, and post-add foam height logged. |

3.3  Hour 6 (T=6) — Feed Initiation Sampling

Sample drawn immediately prior to bolus feed addition; establishes the pre-feed metabolic profile.

| Time Target | Action | Expected Output / Log |
| --- | --- | --- |
| T=6 h | Aseptically draw 10 mL into pre-labelled tube; split into viability and metabolite aliquots. | Sample S-6 logged; aliquots routed to bench + QC. |
| T=6 h +20 min | Initiate bolus feed (Feed-A, 50 mL) via peristaltic pump P-02. | Feed start time, pump RPM, and totaliser reading. |